## Task Analysis
Ref | Task | Subtask | Use Step | Potential Use Error | Hazard | Hazardous Situation | Potential Harm | Severity | Critical Task | Potential Causes | Mitigation / Control | Validation Reference
Ins-exp | Inspect | Confirm AI is not expired |  | User injects with expired drug | Allergic reaction | Reduced therapy | Neurological damage | 4 | yes | Expiry date is not legible, Expiry date is not salient, Not aware that drug can expire, Does not check expiry date | Req # XXX - Font size, color, and location Req # XXX - IFU warning | Validation Study # 1 - Task 1
Ins-dam |  | Check AI is not compromised |  | User injects with damaged AI | Pathogens | Potential Pathogen Exposure | Infection | 4 | yes
Ins-lea |  |  |  | User injects with leaking AI | Allergic reaction | Insufficient drug injected | Death | 5 | yes
Pre-Cap | Prepare | Remove AI from tube | Flip open cap | User does not flip open cap | Allergic reaction | Delay in therapy | Neurological damage | 4 | yes | Does not see flip cap, Flip cap is too difficult to open | Req #XXX - Cap is colored bright yellow, Cap is designed to opened by 5th percentile female child hand strength (Req #YYY
Pre-Tip |  |  | Tip and slide AI out of carrier tube | User does not remove AI out of carrier tube | Allergic reaction | Delay in therapy | Neurological damage | 4 | yes
Pre-001 |  |  | Confirm correct AI | User injects wrong concentration (peds into Adult) | Allergic reaction | Insufficient drug injected | Death | 5 | yes
Pre-002 |  |  |  | User injects wrong concentration (Adult into peds) | Epinephrine | Overdose | Tachycardia | 3 | yes
|  | Orient AI | Hold AI in Fist with Orange Tip pointing downward | User orients AI incorrectly | Needle | Needle Stick into hand & Drug not injected | Bruising | 2 | no
|  | Remove safety release | Pull safety release straight up, without bending or twisting | User does not remove safety release | Allergic reaction | Delay in Therapy | Neurological damage | 4 | yes
|  |  |  | User damages safety release | Allergic reaction | Drug not injected | Death | 5 | yes
| Administer | Position AI | Place orange tip at right angle against thigh | User injects under skin instead of in muscle | Allergic reaction | Insufficient drug injected | Death | 5 | yes
|  |  |  | User injects in location besides thigh | Allergic reaction | Insufficient drug injected | Death | 5 | yes
|  | Inject | Push down AI until click | User does not inject | Allergic reaction | Drug not injected | Death | 5 | yes
|  |  | Hold in place for at least 3 seconds | User does not hold AI for 3 seconds | Allergic reaction | Insufficient drug injected | Neurological damage | 4 | yes
|  |  |  | User moves AI during injection | Needle | Needle injury | Laceration | 2 | no
|  | Remove | Remove AI from leg | User does not remove AI from inject site | Needle | Needle stick injury | Bruising | 2 | no
|  | Massage | Massage injection area for 10 seconds | User does not massage area | Allergic reaction | Slight delay in therapy | No clinical impact | 1 | no
|  |  |  | User does not massage area for at least 10 seconds | Allergic reaction | Slight delay in therapy | No clinical impact | 1 | no
| Get EM help |  | Outside of product scope
| Dispose | Dispose in sharps container | Place used AI in sharps container | User does not dispose of AI in sharps container |  | Needle stick injury | Transmission of blood-borne pathogens | 4
Note on E6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Could potentially roll up into one use error with the previous one: User attempts to inject with AI in carrier tube
Note on E45: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: If a user enter’s an email address that is valid but not theirs, can they recover?
Note on E88: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Need to align with whats in R2/Squire